## Slide 1
ABLATION WITH HIGH POWER SHORT DURATION
Doc. MUDr. Petr Peichl, PhD

## Slide 2
 | Nemám konflikt zájmů | Mám konflikt zájmů | Specifikace konfliktu (vyjmenujte subjekty, firmy či instituce, se kterými Vaše spolupráce může vést ke konfliktu zájmů)
Zaměstnanecký poměr | X |  | 
Vlastník / akcionář | X |  | 
Konzultant | X |  | 
Přednášková činnost |  | x | 
Člen poradních sborů (advisory boards) | X |  | 
Podpora výzkumu / granty | X |  | 
Jiné honoráře (např. za klinické studie či registry) | X |  | 

## Slide 3: What are „conventional“ ablation settings for AF?
25-35W
Low power – long duration (LPLD)
Ablation above conventional
setting may be considered
as high power short duration (HPSD)

## Slide 4
Navistar Thermocool Catheter
Comparison of 40W/30sec vs high power 50-80W 5sec in vitro and in vivo
Compared with 40W/30sec, 50 and 60W for 5 sec achieved transmurality and was safer
No complications for short-duration ablations of 50W-60W
High-Power Short-Duration using conventional cath
Bhashkaran A Europace 2016

## Slide 5: High-Power Short-Duration

## Slide 6: Similar volume, but different shape… HPSD vs LPLD
Bourier JCE 2018

## Slide 7: High-Power Short-Duration using conventional cath
51pts undergoing AF ablation
50W using Tacticath
  Length guided by loss of capture (mean of 11±4sec)
Procedure time
  101±20min
RF time
  14.9±4.3 min
No complications
AF free at one year 86%
Winkle RA, JICE 2017

## Slide 8: HPSD vs LPLD using conventional catheters
Winkel JCE 2021
Studies evaluated:
CF, but also nonCF cath
Power ranges from 45W/20sec
50W/10sec
70W/5-7sec
Some studies used AI
Notes: Conventional catheters 50W-70W up to 7 sec.

## Slide 9: HPSD vs LPLD using conventional catheters
HPSD is faster and less RF is required
Winkel JCE 2021
Notes: Conventional catheters 50W-70W up to 7 sec.

## Slide 10: Outcome and complications HPSD vs LPLD using conventional catheters
Better outcome in some studies
Safety profile is similar to conventional ablation setting…
Winkel JCE 2021

## Slide 11: High power requires better control

## Slide 12: High power may have serious consequences Lessons from animal work
ICE during „pop“ in LV apex
Gross anatomy at „pop“ site

## Slide 13: To prevent tissue overheating, temperature should be measured during ablation
Measurement of temperature at the catheter-tissue interface is limited by the confounding effect of the cold irrigation fluid during ablation
Nonirrigated tip - Power 25W, T 55⁰C
Irrigated tip - Power 35W, T 43⁰C
Notes: Novel catheter design may overcome this limitation and provide better surrogate for lesion size control

## Slide 14: QDOT – Biosense Webster Catheter design
Improved temperature sensing by 6 thermocouples embedded in the circumference of the tip electrode
  Distal 75µm from tip
  Proximal 3mm from tip
Improved irrigation system allowing adequate irrigation at low rates
Three microelectrodes
  Providing high-resolution electrograms
  Allow display of discrete local signals

## Slide 15: vHPSD ablation using QDOT cath Bench study
Leshem E;JACC Clinical Electrophysiology 2018

## Slide 16: Comparison of vHPSD with/without temperature control
Ablation in beating heart
No popping with temperature limit to 65ºC
Three pops occurred (1.7%) without temperature limit
Mean temperature during pops >85ºC
Leshem E;JACC Clinical Electrophysiology 2018

## Slide 17: QDOT – Biosense Webster Workflow
RF energy modulation to reach the selected target temperature (47⁰C, max 50⁰C)
QMODE Irrigation titration:
  Irrigation rate and power vary automatically between 4ml and 15ml based on tip temperature for optimized power delivery
  Power <35W irrigation 4ml/min, when temp >47⁰C irrigation increased to 15ml/min
  Power >35W irrigation 15ml, when temp <42⁰ irrigation decreased to 4ml/min
QMODE +:
  Ablation with 90W 4sec – True vHPSD
Notes: Irrigation rate and power vary automatically between 4ml and 15ml based on tip temperature for optimized power delivery
Ensure the tip is within the allowed temperature range to avoid over-heating or over-cooling
Ablation Index standardizes the PAF procedural Workflow and allows reproducibility across multiple operators.*

## Slide 18: vHSPD vs conventional ablation LPLD
Leshem E;JACC Clinical Electrophysiology 2018

## Slide 19: QDOT FAST Multicenter trial
52pts with PVI
90W/4sec
All veins isolated with study protocol
Procedure time 105min
One AE ulcer and hemorrhage at endoscopy
Six pts (11%) had silent cerebral lesions
Reddy V. JACC EP 2019

## Slide 20: Other novel catheter tips for HPSD

## Slide 21: Diamond tip catheter (Medtronic)
7.5F externally irrigated (6 ports) catheter with 4.1mm tip electrode
Composite tip providing higher resolution electrograms
Tip is embedded with 2 industrial-grade diamonds allow rapid heat shunting
  (2 orders better thermal diffusivity than Ir-Pt)
Six thermocouples at the tip
Dedicated RF generator (ACT)
Iwasawa J et al. JACC 2017;70:542-53
Notes: Tip is made of iridium –platina
Composite tip – during ablation works as one

## Slide 22: Clinical experience with Diamond tip
482pts with pAF
Randomized to PVI by:
  Diamond tip
    50W 60ºC
  Tacticath abl
    30W
Better safety profile
Shorter RF time
Noninferiority outcome

## Slide 23: Affera Inc.
Sphere-9 is an 8F steerable ablation catheter
RF energy is delivered through a 9mm expandable irrigated tip
9 uniformly distributed temperature sensors and mini-electrodes enable high density mapping and temperature controlled energy delivery
The system also includes a proprietary ablation generator, pump and mapping system
  Radiofrequency High power– short duration lesions
    3-5 sec ablations with ~450W
  Pulse field (electroporation)
Barkagan Circ EP 2019

## Slide 24: Ablation using lattice tip with HPSD provided durable lesions
Reddy..Peichl..Kautzner JACC EP 2020

## Slide 25: High-Power Short-Duration Summary
Advantages
  Faster
  Wider lesions
  Superficial (lower risk for collateral damage)
  As safe as conventional ablation
Disadvantages
  Not suitable for thick tissue
Notes: Clearly better?
Lesion index validated for 35W
Steerable sheat does influence?
Same power ant/post
Not deep – but cumulative energy